GreenSeeker™ (Red NDVI) Performance Review

Testing Summary (EXTERNAL)

Objective:

- To determine performance characteristics of GreenSeeker™ Rev. G (Red NDVI) sensors

Testing Parameters:

- All tests were conducted under steady state conditions unless otherwise noted.  At least two minutes were allowed for the sensor to reach steady state before testing commenced.

- Step Test  - dynamic response

- Target description:  Teal and white strip target

- Height held constant at 1 m from target to photodiode window

- Target moved from teal strip to white strip to teal strip, velocity = 0.08m/sec

- NDVI changes as sensor moves from teal area to white area, then to teal area again

- Transition zone distance = 2.7 cm

0

0.2

0.4

0.6

0.8

1

0

0.1

0.2

0.3

0.4

0.5

0.6

Transition

Area

Transition Area

Teal Strip

Teal Strip

White Strip

Distance Traveled, m

NDVI

0.7

Transition Zones

0.8

0.7

0.8

# Figure 1. Dynamic Response of GreenSeeker Sensor

- Conclusion:  Sensor responded appropriately with little ramping effect due to target changes.

- Time Effect on Sensor Signal

- Sensor was mounted stationary over a teal and white strip target

- Height was held constant at 1 m from target to photodiode window

- Test was conducted for two minutes

# Figure 2. Two minute elapsed time test of GreenSeeker sensor

- Average NDVI = 0.816, standard deviation in NDVI = 0.0056

- Conclusion:  Sensor variation in NDVI throughout the two-minute test was small.

- Temperature Test

- Tested at 7°C, 21°C, and 42°C

- At least two minutes allowed for sensor to reach steady state at each temperature

- Three heights (0.94 m, 1 m, and 1.3 m) tested at each temperature

- Teal strip target

| Height | NDVI at Different Temperatures |  |  | Temperature Coefficient, Change in NDVI per °C |
| --- | --- | --- | --- | --- |
|  | 7°C | 21°C | 42°C |  |
| 0.94 m | 0.804 | 0.814 | 0.799 | 0.0004 |
| 1.0 m | 0.815 | 0.817 | 0.8003 |  |
| 1.3 m | 0.811 | 0.819 | 0.7938 |  |

# Table 1. Variation due to temperature of GreenSeeker sensor

- Results:  Temperature has little effect on NDVI within the specified temperature range.

- Sensor-to-target Distance Test

- Tested at 21°C

- Teal strip target 0.91 m wide

NDVI = .00003height + 0.8296

R2 = 0.0542

(equation applies to these experimental conditions)

Std. Deviation = 0.005858 NDVI

## Figure 3. NDVI change with change in GreenSeeker height

- Conclusion:  NDVI is independent of height over the range tested (Slope = 0.0003, R2 = 0.0542)

- Standard deviation at each height is shown in Figure 4

Std. Dev. = 0.0037e0.4341distance

R2 = 0.08

## Figure 4. Standard deviation of NDVI with respect to change in sensor height

- Conclusion:  Variability of data increased slightly as height increased.  However, the average NDVI was approximately the same.

- Sensitivity across field of view

- Tested 5 cm white plate on teal target and 5 cm teal plate on white target

- 5 cm plates moved across target strip in 2.54 cm increments

- Height held constant at 0.9 m from target to photodiode window

- Tested at 21°C

- Tested over the designed field of view of the WeedSeeker/GreenSeeker system

Figure 5. Sensitivity to white plate on teal strip background across GreenSeeker sensor field of view

### Figure 6.  Sensitivity to teal plate on white strip background across GreenSeeker sensor field of view

- Conclusions:  NDVI varied less than 8% across the field of view in the white plate on teal background test and 12% in the teal plate on white background test.

- Sunlight effect

- Height was held constant at 1.0 m from target to photodiode window

- Various turf targets were used

- Data was gathered in direct sunlight and with sunlight blocked with black 3 mil plastic cover

| Test Number | With Sunlight (NDVI) | Blocked Sunlight (NDVI) | Change in NDVI |
| --- | --- | --- | --- |
| 1 | 0.8104 | 0.8106 | -0.0002 |
| 2 | 0.5850 | 0.5970 | -0.0120 |
| 3 | 0.8126 | 0.8119 | 0.0007 |
| 4 | 0.8125 | 0.8022 | 0.0103 |
| 5 | 0.8006 | 0.7988 | 0.0018 |
| Mean NDVI Difference = 0.00012 |  |  |  |
| Standard Deviation of NDVI change = 0.00795 |  |  |  |

Table 2.  Effect of sunlight on GreenSeeker sensor

- Conclusion:  Sunlight had little effect on NDVI (Mean difference = 0.00012 NDVI)

- Plant-on-soil height test

- Tested at 25°C, outdoors

- Sensor positioned directly above a single plant (approx. 15 cm diameter) with a clear soil background.

- Clear soil NDVI = 0.100

NDVI = -0.0606Distance + 0.4675

r2 = 0.41

(equation applies to these experimental conditions)

std. Deviation = 0.0121

## Figure 3. NDVI change with change in GreenSeeker height

- Conclusion:  Sensor can detect foliage with little variation in NDVI over a range of distances, 0.81 m to 1.22 m (32 in. to 48 in.) from canopy to photodiode window.